Review Criteria for UV Special Section in Journal of Optical Engineering

The manuscript that you are asked to review is entitled:  UV aerosol absorption experiment (2002-2004): 1. UV-MFRSR calibration and performance at GSFC.

Please review this manuscript in both of its scientific merit and composition quality.

Authors: N.A. Krotkov, P.K. Bhartia, J.R. Herman, J.R. Slusser, G. Scott, G. Labow, T.F. Eck, and B.N. Holben

MS#: UV10

Please Return by: June 21, 2004

| SCIENTIFIC EVALUATION | YES | MARG | NO | COMMENTS |
| --- | --- | --- | --- | --- |
| Subject dealt with the most recent advances in the field | x |  |  |  |
| Data and/or results have international value | x |  |  |  |
| Theory and/or methods are new and original | x |  |  |  |
| Conclusions are correctly drawn | x |  |  |  |
| Scientific literature is used appropriately | x |  |  |  |

Use separate sheet(s) for detailed comments.  Please address the following questions:

(1)  What are the important contributions of the paper?

(2)  What is your evaluation of the technical aspects of the paper?

Return the manuscript only if you have made comments on it that you wish to be communicated to the author(s).

Please comment on the following aspects of PRESENTATION.

1)  Does the title properly reflect the content of the paper?

2)  Is the Abstract informative and does it accurately convey the findings?

3)  Is the paper organized in a logical manner?

4)  Please point out specifically any areas or passages that are unclear.

5)  Is the paper of appropriate length?

6)  Are the figures a) necessary; b) informative; c) of sufficient standard?

7)  Is the manuscript clearly presented in standard scientific English?

RECOMMENDATION:

____Accept  ____x_ Accept-Minor revision ______ Major revision _______ Reject

Do you wish to review the manuscript after revision?   _______Yes  _____x_  No

Review of "UV aerosol absorption experiment (2002-04): 1. UV-MFRSR

calibration and performance at GSFC" by Krotkov et al.

The paper presents very interesting new ideas for instrument calibration. The

procedure is properly described (with some exceptions listed below) and the

results are illustrated in a clear and straightforward way. I recommend

publication of the manuscript after consideration of the following minor

points.

Comments to the specific points to be addressed according to the review

guidelines:

1) Yes, the title properly reflects the contents of the paper

2) Yes, the abstract is informative and conveys the findings

3) Yes, the paper is organized in a logical manner

4) Some small points are unclear, see specific points below

5) Yes, the paper is of appropriate length

6) The figures are necessary and informative, but the quality of

some is poor, at least in the version which I received

electronically; are these jpeg images? If yes I would suggest

a different format with lossless compression (gif, eps, ...)

7) Yes, the manuscript is clearly presented in standard scientific

English.

Specific points:

page 3, 1st sentence: The formulation is a bit unclear: harmful UV radiation is not a consequence of tropospheric chemistry

page 3, end of 2nd paragraph, "reductions in soot emissions ...

may make smog problems worse". Although correct, this is a dangerous

statement which could easily be abused as an excuse for not reducing

emissions; I would suggest to formulate this more carefully

page 4, second line in "2. Instrumentation": "currently operates" instead of "currently operate"

page 6, 3rd paragraph: "dark current" is not a good term for a voltage

page 7, last line: It is not clear why the "dark current" is added to the total irradiance

page 8, 1st line after equation (1): f_R is NOT given in equation (1); it is probably the cosine correction of the direct voltage but that isn't stated explicitely

page 8, first paragraph of 3.3: Symbols V_n0 and V_0 are used for the same quantity; only one should be used consistently throughout the paper

page 8, point 1)  the same symbol P is used for relative and absolute

pressure?

page 8, point 5)  "either diffuse or total" instead of "either diffuse of total"

page 9, caption of Figure 5 and other places: pressure is given in atm which is not a SI unit; SI units should be used throughout the paper (hPa, or at least mbar)

page 9, caption of Figure 5: Was the ozone column really 500 DU? If yes, why?

page 9, caption of Figure 5: The plot shows the spectral irradiance MULTIPLIED with the filter function, not CONVOLVED

page 9, caption of Figure 5: lambda_dir and lambda_neff are use here;

both quantities are not explaiend. "neff" occurs a some places while "eff" is used at others, probably for the same reason as above with V_n0 where the n denotes the channel number which is omitted at other places; lambda_dir is not explained - it is probably the same as lambda_rad?

Please use consistent symbols!

page 10, after equation (3): at a first glance it is not clear how two

quantities (V0 and lamdba_rad) can be derived from only one equation (2), but the simple explanation is that (2) is in fact two equations; I suggest to split it in two parts to facilitate understanding

page 16: <ln (r**2 V0)> is not explained

page 17: I had a deja vu here because the text is nearly exactly identical to page 14; one is really tempted to check if one has been transported back a few pages.

Conclusions: Although the calibration method shows very good performance, one additional uncertainty remains in the absolute calibration: the uncertainty of the extraterrestrial spectrum. Could the authors comment on that?